Hydrogen Sulphide Sensor Specification

MICROceLTM HS - Hydrogen Sulphide

Performance Characteristics

Nominal Range

Maximum Overload

Expected Operating Life

Output Signal

Resolution

Temperature Range

Pressure Range

T50 Response Time

T90 Response Time

Relative Humidity Range

Typical Baseline Range

(pure air)

Maximum Zero Shift

(+20°C to +40°C)

Long Term Output Drift

Recommended Load

Resistor

Bias Voltage

Repeatability

Output Linearity

0-100ppm

1000ppm

Two years in Air

0.105 ± 0.025

0.2ppm

-40°C to +50°C

Atmospheric ± 10%

<10 seconds

25 to 40 seconds typically

15 to 90% non-condensing

-2 to +1ppm equivalent

<0.2ppm equivalent

<5% signal loss/year

10Ω

Not required

<2% of signal

Linear

Outline Dimensions

17 ±0.2

5 CRS

Typ

2.8 ±0.2

Typ

 8.25 ±0.2

Product Label

N.B. All performance data is based on conditions at 20°C,

50%RH, and 1013mBar unless otherwise noted.

Physical Characteristics

Weight 1.2g (approx.)

All dimensions in mm

All tolerances ±0.15mm unless otherwise stated

Position Sensitivity

Storage Life

Recommended Storage

Temperature

Warranty Period

None

Six months in CTL container

0°C to 20°C

12 months from date of

despatch

Doc. Ref.: microcelhs.pmd Issue 1.0

11th March 2004

Doc. Ref.: microcelhs Rev 02 Issue 2 ECN I 1281

2nd January 2007

Page 1 of 2

Page 1 of 2

17±0.2

4±0.2

2.75±0.2

9.75±0.2

1.85±0.1

Hydrogen Sulphide Sensor Specification

MICROceL™ H2S Sensor - Output vs Temperature

115

110

105

100

95

90

85

80

Mean of 15 Sensors

75

70

Bars denote ±1

-40

-30

-20

-10

0

10

20

30

40

50

Temperature (°C)

Effect of Temperature on the Shift in Mean Air Baseline Signals of Microcel HS Sensors

0.5

0.4

0.3

0.2

0.1

0.0

-0.1

-0.2

-0.3

-0.4

-0.5

-50

-40

-30

-20

-10

0

10

20

30

40

50

60

Temperature / °C

Cross-sensitivity Data

CiTiceLs may exhibit a response to certain gases in a sample other than the target gas. MICROceL™H2Ss have been tested

with a number of commonly cross-interfering gases and the results are given below. The table shows the typical response to be

expected from a sensor when exposed to a given test gas concentration (relevant to safety, e.g. TLV levels).

Gas

Carbon Monoxide:

Sulphur dioxide:

Conc.

15ppm

5ppm

MICROceL™H2S

0.1ppm

~1ppm

Gas

Hydrogen:

Nitrogen Dioxide:

Conc.

10000ppm

5ppm

MICROceL™H2S

<10ppm

~ -1ppm

Nitric oxide:

35ppm

~0.7ppm

**For details of other possible cross-interfering gases contact City Technology.**

Every effort has been made to ensure the accuracy of this document at the time of printing. In accordance with the company’s policy of continued product improvement City

Technology Limited reserves the right to make product changes without notice. No liability is accepted for any consequential losses, injury or damage resulting from the use

of this document or from any omissions or errors herein. The data is given for guidance only. It does not constitute a specification or an offer for sale. The products are always

subject to a programme of improvement and testing which may result in some changes in the characteristics quoted. As the products may be used by the client in circumstances

beyond the knowledge and control of City Technology Limited, we cannot give any warranty as to the relevance of these particulars to an application. It is the clients’ responsibility

to carry out the necessary tests to determine the usefulness of the products and to ensure their safety of operation in a particular application.

Performance characteristics on this data sheet outline the performance of newly supplied sensors. Output signal can drift below the lower limit over

Doc. Ref.: microcelhs.pmd Issue 1.0                Page 2 of 2

11th March 2004

Doc. Ref.: microcelhs Rev 02 Issue 2 ECN I 1281

2nd January 2007

Page 2 of 2

Output(%ofsignal@20°C)

BaselineShift/ppmH2S

time.

Thank you for evaluating AdreamSoft PDF to Word

You can only convert 3 pages with the trial version

To get all the pages converted, you need to purchase the software from

http://www.allpdftools.com/purchase/buy-pdf-to-word.html